# NOTICE OF CLEANING AND/OR DECONTAMINATION

Complete this form for each individual item to certify it has been cleaned, decontaminated, and contains no hazardous materials prior to service, surplus, moving, or handling by UW Facilities personnel or contracted vendors. The form must be completed by a UW employee knowledgeable of the hazardous materials used in, on and/or around the item. Complete all sections.

UW Facilities staff and vendors have the right to refuse service if the item, or the surrounding area, is not visibly clean, appropriately decontaminated, or safe. Any items missing a form will not be moved and will require rescheduling.

## ITEM INFORMATION

Specify type of equipment, appliance, surface, or furniture: ______________________________

Make/Model (if applicable): ______________________________ UW Asset tag# (if applicable): _____________________

## REASON for CLEANING/DECONTAMINATION (SELECT one)

Service on site	  UW Surplus 	  Moving to (building/room):________________________________________

Disposal at (location):______________________________   Other (please specify): _____________________________________

## HAZARD EXPOSURE & DECONTAMINATION (SELECT ALL THAT APPLY)

*Biological Safety cabinets require additional notice of gaseous paraformaldehyde decontamination posted by EH&S prior to servicing or removal from the room.

To the best of my knowledge, the item described above may have been exposed to and managed as follows:

Radioactive materials: item was emptied and decontaminated according to EH&S instructions, and final clearance was obtained from EH&S Radiation Safety (required).

Biohazards: item was emptied and decontaminated according to EH&S instructions.

Hazardous chemicals (poisons, toxics, corrosives, solvents): item was emptied and decontaminated according to EH&S instructions. Note: UW Surplus does not accept chemical containers.

None (no exposure): item was never exposed to hazardous materials and was washed with warm, soapy water.

- CLEANING & PREPARATION

All internal and external surfaces have been wiped down and cleaned to remove visible residue, encrusted material, and dirt including the top, sides, back, inside, and hard-to-reach areas behind or underneath the item.

If applicable:

- Remove or deface hazard warning labels/signs (e.g., biohazard, radioactive, etc.).

- Remove all contents, fluids, internal and external tubing, and filters (except biological safety cabinet filters)

- CERTIFICATION

I certify that all steps on this form have been completed and the item is clean and safe to handle without PPE. I understand cleaning must be finished prior to pickup or service and I am the point of contact for any questions.

Name/Title:							Phone:

Department: 							Building/Room:

Signature: 							Date:

GUIDELINES FOR CLEANING AND DECONTAMINATION

- Review UW Surplus preparation guidelines, including lists of prohibited and special consideration items, before submitting any items for surplus.

- ALL EQUIPMENT must be cleaned inside and out of all visible residue and encrusted material.

- Remove or deface hazard warning labels and signage (e.g., biohazard, radioactive, etc.).

- Remove all internal and external tubing and filters (does not apply to biological safety cabinets filters).

- For biological safety cabinets being moved or serviced, contact EH&S Biological Safety Cabinet program at 206.542.9510 or labair@uw.edu to schedule decontamination at least two weeks prior to the move date.

- For fume hoods being serviced, remove all obstructions from around the area, remove all containers inside the fume hood, and thoroughly wash interior surfaces with warm, soapy water. All residues must be cleaned from walls, floor, and sash of the hood.

## RADIOACTIVE MATERIALS decontamination

- A thorough radiation survey of all accessible surfaces must be performed with an appropriate instrument.

- If radioactive contamination is detected, the equipment must be cleaned with small amounts of warm detergent water to avoid splash. Blot dry with paper towels. Commercial radiation decontamination solutions containing chelating agents may be helpful.

- Resurvey to ensure contamination has been removed to less than 100 counts per minute per 100 cm2.

- Contact EH&S Radiation Safety at radsaf@uw.edu / 206.543.0463 for final clearance of the item.

## BIOHAZARD decontamination

- Remove biological material from the equipment.

- Clean the equipment with warm, soapy water and scrub as necessary.

- Sanitize with a 1:10 bleach solution. After 10 minutes contact time, rinse metal surfaces as bleach is corrosive.

## HAZARDOUS CHEMICALS decontamination

- Remove, drain, or discharge chemicals from the equipment. Collect for reuse or hazardous waste disposal.

- If applicable, use an inert gas or liquid to purge or rinse out chemical residues.

- Any rinsate may need to be disposed of as hazardous waste as well. Email chmwaste@uw.edu or call 206.616.5835 if you have questions about chemical or rinsate disposal.

- Clean the equipment with warm, soapy water and scrub as necessary.

Visit the Disposal, Recycling, and Surplus webpage for more details.

Contact Environmental Health & Safety at ehshelp@uw.edu or 206.543.7262.